# Base Information

| Objective | The aim of this specification document is to verify the implementation of the Straumann IOS software design specifications outlined in DEV-0044563 STMN IOS SDD - Scan Workflows [1] |
| --- | --- |
| Version No. | 6 |
| PRS [2] Requirements tested | 6.001 |

# Related Test Report

| Filename | DEV-0044595 STMN IOS Scan Workflows Verification Report [3] |
| --- | --- |

# Final Result

A summary evaluation of all results in the related test report is given by the independent reviewer on the front page of the related test report. The final result is stated at the end of the quality control.

- Pass, if everything in the report has been passed.

- Fail, if one or more test steps failed.

The reviewer must give their full name (first name or initial and full last name), the date of the review and signature to the report.

# Model-Specific Test Cases

All test cases are applicable to all scanner models, except those explicitly listed below as specific to certain models.

| Scanner Model | Specific Test Case(s) | Remarks |
| --- | --- | --- |
| Model ST 200E | n/a | This test case is specific to Model ST 200E |
| Model ST 300E | n/a | This test case is specific to Model ST 300E |

# Test Specification

## TC06101 – Basic workflow – perform 3 basic scans

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | - Set the User “Scan Preferences” for first arch to scan is to “Always the lower arch” (under Options) - Create a new case with no additional scans selected. - User is in the Treatment Plan interface. | n/a | n/a | n/a |
| 1 | Proceed to the next step in the case’s workflow using the navigation button. | n/a | The user can enter the Scan interface using the navigation button. The first scan is the ‘Lower’ arch scan The Scan interface displays the scan of the ‘Lower’ arch first, then the ‘Upper’ scan, then the ‘Bite’ scan. | Preparation |
| 2 | Perform the first scan available. | 6881_101 | The scan of the Lower Arch can be performed successfully. | Observation / Result documentation |
| 3 | Open the HD areas toolkit and add one over a tooth. Go to the next scan step. Then, go back to the first scan. | 6881_308 | HD marker is added over the tooth. After going to the second step and back to the first, the cylinder previously added was saved and it is visible. | Observation / Screenshot |
| 4 | Repeat steps 2 and 3 for every step available in the workflow, until the last scan is completed. Count the total number of scans performed. | 6881_101 | A total of 3 scans is performed: Lower, Upper and Bite. | Observation / Result documentation |

## TC06118 – Simple Scan workflow – perform 1 to 3 basic scans

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | - Set the User “Scan Preferences” for first arch to scan is to “Always the lower arch” (under Options) - Create a new Simple Scan case - User is in the Treatment Plan interface. | n/a | n/a | n/a |
| 1 | Select Lower option and proceed to the next step in the case’s workflow using the navigation button. | 6881_101 | The Scan interface displays the Lower single scan step. | Observation / Screenshot |
| 2 | Perform the scan. | 6881_101 | The scan of the Lower is performed successfully. | Observation / Result documentation |
| 3 | Proceed to the Review step to perform post-processing. | n/a | n/a | Preparation |
| 4 | Go back to Treatment Plan, select both arches and proceed to the next step in the case’s workflow using the navigation button. | 6881_101 | The user enters the Scan interface using the navigation button. The Scan interface displays the scan steps of the ‘Lower’ arch first, then the ‘Upper’ scan, then the ‘Bite’ scan. The review step is disabled. | Observation / Screenshot |
| 5 | Perform all the scans. | 6881_101 | The scan of the Lower, Upper and Bite is performed successfully. | Observation / Result documentation |
| 6 | Open the HD areas toolkit and add one over a tooth. Go to the next scan step. Then, go back to the first scan. | 6881_308 | HD marker is added over the tooth. After going to the second step and back to the first, the cylinder previously added was saved and it is visible. | Observation/ Screenshot |

## TC06102 – Complex workflow – Crown and Bridge –  HD scanning approach information

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create two new users (“User_1” and “User_2”) Execute the steps from 1 to 6 with an Administrator too | n/a | n/a | n/a |
| 1 | Create a Crown and Bridge case. | 6881_109 | In Treatment Plan, a pop-up is opened to remind the user to add the HD markers on preparations and scanbodies. | Observation/ Screenshot |
| 2 | Move to the Home. | 6881_109 | Home is displayed. | Observation/ Result documentation |
| 3 | Reopen the case previously created and access the Treatment Plan. | 6881_109 | The pop-up is still displayed. | Observation/ Screenshot |
| 4 | Mark the option to not show the pop-up again and close the pop-up. | 6881_110 | The pop-up is closed. | Observation/ Screenshot |
| 5 | Move to the Home and then, reopen the same case. | 6881_110 | In Treatment Plan, the pop-up is not displayed anymore. | Observation/ Result documentation |
| 6 | Create a new Crown and Bridge case. | 6881_110 | In Treatment Plan, the pop-up is not displayed anymore. | Observation/ Result documentation |
| 7 | Login with a different user and create a Crown and Bridge case. | 6881_109 | In Treatment Plan, the HD reminder pop-up is displayed. | Observation/ Screenshot |

## TC06103 – Basic and Complex workflow – HD areas toolkit availability

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | n/a | n/a | n/a | n/a |
| 1 | Create a ClearCorrect case. Enter the Scan step. | 6881_301 | The HD areas toolkit is not displayed. | Observation/ Screenshot |
| 2 | Create a Simple Scan case, both arches. Enter the first arch. | 6881_301 | The HD areas toolkit is displayed. | Observation/ Screenshot |
| 3 | Scan the first arch. | 6881_301 | The HD areas toolkit is enabled when the scanning is stopped. | Observation/ Screenshot |
| 4 | Enter the second arch scan view. | 6881_301 | The HD areas toolkit is disabled. | Observation/ Screenshot |
| 5 | Scan the second arch. | 6881_301 | The HD areas toolkit is available when the scanning is stopped. | Observation/ Screenshot |
| 6 | Enter the bite scan view and proceed with the scan. | n/a | n/a | Preparation |
| 7 | Create a Crown and Bridge case, with pre-op. Perform all the scans. | 6881_301 | The HD areas toolkit is displayed and available after scanning except in pre-op. | Observation/ Screenshot |
| 8 | Create a Full Arch case, with pre-op. Perform all the scans. | 6881_301 | The HD areas toolkit is not displayed. | Observation/ Screenshot |

## TC06107 – Basic workflow – Guidance Tooltip

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create two new users (“User_1” and (“User_2”) Log in as “User_1” Create a new case with only the 3 basic scans. | n/a | n/a | n/a |
| 1 | Enter the Lower and Upper scan steps for the first time | 6881_106 6881_107 | A guidance information is displayed to help the user to scan and mark the Lower and Upper arches. | Observation / Screenshot |
| 2 | Close the guidance information in Lower scan step. Check the guidance in Upper scan step. | 6881_107 6881_108 | The guidance information is closed for both Lower and Upper scan steps. | Observation / Screenshot |
| 3 | Navigate out of the Lower and Upper scan steps, then enter again each step. | 6881_108 | The guidance information remains closed in the Lower and Upper scan step upon re-entry. | Observation/ Result documentation |
| 4 | Perform the Lower and Upper scans. | 6881_108 | The guidance information remains closed in the Lower and Upper scan steps during and after scanning. | Observation/ Result documentation |
| 5 | Open a new case with the current user and enter the Lower and Upper steps. | 6881_108 | The guidance information remains closed in the Lower and Upper scan steps. | Observation/ Result documentation |
| 6 | Open an existing case with the current user and enter the Lower and Upper scan steps. | 6881_108 | The guidance information remains closed in the Lower and Upper scan steps. | Observation/ Result documentation |
| 7 | Shutdown the software, then launch it again. Log in as “User_1”. Open a new case and enter the Lower and Upper scan steps. | 6881_108 | The guidance information remains closed in the Lower and Upper scan steps. | Observation/ Result documentation |
| 8 | Log out and log in as “User_2” Create a new case and enter the Lower and Upper scan steps. | 6881_108 | Guidance information is displayed (opened) to help the user to scan and mark the Lower and Upper arches. | Observation/ Result documentation |
| 9 | Close the guidance information, then open it again. | 6881_107 | The user can collapse and open again the guidance information. | Observation/ Result documentation |
| 10 | Create another case with “User_2”. Enter the Lower and Upper scan steps and perform the scan. | 6881_107 | The guidance information is displayed when the user opens another case. It remains displayed during the Lower and Upper scanning. | Observation/ Result documentation |
| 11 | Log out and Login as the previous user (“User_1”). Reopen any of the cases created in this test for this user and enter the Lower and Upper scan steps. | 6881_108 | The guidance information remains closed in the Lower and Upper scan steps. | Observation/ Result documentation |

## TC06108 – Complex workflow – Guidance Tooltip – Crown case

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create two new users (“User_1” and “User_2”) Log in as “User_1” Create a new tooth borne restoration case with additional “Lower pre-op” scan | n/a | n/a | n/a |
| 1 | Enter the Lower pre-op scan step. | 6881_106 6881_107 | Guidance information is displayed to help the user to scan the arch. | Observation / Screenshot |
| 2 | Close the guidance information in Lower pre-op and perform the scan. | 6881_107 | The guidance information remains closed during the scan. | Observation/ Result documentation |
| 3 | Enter the Upper scan step. | 6881_106 6881_107 | Guidance information is displayed to help the user to scan and add HD areas to the Upper arch. | Observation / Screenshot |
| 4 | Close the guidance information in the Upper scan step. | 6881_107 | The guidance information is closed. | Observation/ Result documentation |
| 5 | Enter the Lower scan step. | 6881_106 6881_107 | Guidance information is displayed to help the user to scan the preparations and the emergence profiles on the Lower arch plus the HD areas. | Observation / Screenshot |
| 6 | Close the guidance information in the Lower scan step. | 6881_107 | The guidance information is closed. | Observation/ Result documentation |
| 7 | Shutdown the software, then launch it again. Log in as “User_1”. Open a new case with the same treatment plan configuration as seen in the Setup step. Enter all the scan steps. | 6881_108 | The guidance information remains closed in the Lower pre-op, Upper and Lower scan steps. | Observation/ Result documentation |
| 8 | Log out and log in as User_2 Create a new case with the same treatment plan configuration as seen in the Setup step. Enter all the scan steps. | 6881_108 | The guidance information is displayed in Lower pre-op, Upper and Lower scan steps. | Observation/ Result documentation |

## TC06109 – Complex workflow – Guidance Tooltip – Implant case

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create two new users (“User_1” and “User_2”) Log in as “User_1” Create a new implant case with additional “Lower pre-op”, “Upper emergence profile” and “Lower emergence profile” scans | n/a | n/a | n/a |
| 1 | Enter the Lower pre-op scan step. | 6881_106 | Guidance information is displayed to help the user to scan the arch. | Observation / Screenshot |
| 2 | Close the guidance information in Lower pre-op and perform the scan. | 6881_107 | The guidance information remains closed during the scan. | Observation/ Result documentation |
| 3 | Enter the Upper emergence profile scan step. | 6881_106 6881_107 | Guidance information is displayed to help the user to scan the Upper emergence profile. | Observation / Screenshot |
| 4 | Close the guidance information in Upper emergence profile and perform the scan. | 6881_107 | The guidance information remains closed during the scan. | Observation/ Result documentation |
| 5 | Enter the Lower emergence profile scan step. | 6881_106 6881_107 6881_108 | Guidance information is displayed to help the user to scan the mesh. | Observation/ Screenshot |
| 6 | Close the guidance information in the Lower emergence profile scan step. | 6881_107 | The guidance information is closed. | Observation/ Result documentation |
| 7 | Enter the Lower scanbody scan step and check the guidance information | 6881_108 | Guidance information is displayed to help the user to scan the mesh. | Observation/ Screenshot |
| 8 | Close the guidance information in the Lower scanbody scan step. | 6881_107 | The guidance information is closed. | Observation/ Result documentation |
| 9 | Enter the Upper scanbody scan step and check the guidance information | 6881_108 | The guidance information is closed. | Observation/ Result documentation |
| 10 | Shutdown the software, then launch it again. Log in as “User_1”. Open a new case with the same treatment plan configuration as seen in the Setup step. Enter all the scan steps. | 6881_108 | The guidance information remains closed in the Lower pre-op, Upper emergence profile, Lower emergence profile, Lower scanbody and Upper scanbody scan steps. | Observation/ Result documentation |
| 11 | Log out and log in as User_2 Create a new case with the same treatment plan configuration as seen in the Setup step. Enter all the scan steps. | 6881_108 | The guidance information is displayed in Lower pre-op, Upper emergence profile, Lower emergence profile, Lower scanbody and Upper scanbody scan steps. | Observation/ Result documentation |

## TC06111 – Complex Workflow –Tooth-borne Workflow

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Log in as a regular user (non-admin) | n/a | n/a | n/a |
| 1 | Create a new case and in the Treatment Plan configure the following: - Configure one crown restoration on each arch, i.e. on Upper and Lower - Search and select the Upper pre-op and Lower pre-op additional scans | 6881_102 | “Upper pre-op” and “Lower pre-op” scans are available to select in the Additional Scans box. The Upper pre-op and Lower pre-op additional scans are selected when clicked | Observation/ Screenshot |
| 2 | Click Next and observe the different workflow steps displayed under the Scan steps menu | 6881_102, 6881_201 | The following steps are displayed: Upper pre-op, Lower pre-op, Bite Upper, Lower The two pre-op scan steps are before/above the Bite step, and the rest are after/below the Bite step. Every scan step after/below the Bite is disabled (grayed out) | Observation/ Screenshot |
| 3 | Perform the Lower pre-op scan and when stopping the scan, inspect the Scan steps menu | 6881_102, 6881_201 | The scan is performed successfully When the scan is stopped, the Lower step becomes enabled in the Scan menu | Observation/ Screenshot |
| 4 | Perform the Upper pre-op scan and when stopping the scan, inspect the Scan steps menu | 6881_102, 6881_201 | The scan is performed successfully When the scan is stopped, the Upper step becomes enabled in the Scan menu | Observation/ Screenshot |
| 5 | Click the Lower scan step in the Scan steps menu | 6881_203 | Clicking the Lower step navigates to that step, and: The ancestor mesh acquired in the Lower pre-op step is displayed | Observation/ Screenshot |
| 6 | Open the HD areas toolkit and add one over a tooth. | 6881_301 | HD marker is added over the tooth. | Observation/ Screenshot |
| 7 | Go to the Upper scan step. Then, go back to Lower scan. | 6881_308, 6881_313 | After going to the Upper scan and back to the Lower scan, the HD area previously added was saved and it is visible. | Observation/ Screenshot |
| 8 | Click the Upper scan step | 6881_203 6881_031 | Clicking the Upper step navigates to that step, and: - The ancestor mesh acquired in the Upper pre-op step is displayed, as well as the HD area | Observation/ Screenshot |
| 9 | Scan in Upper. Then, remove the scan. | 6881_206 | Only the new scan is removed. The derived scan is kept. | Observation/ Result Documentation |

## TC06112 – Complex workflow – Implant-borne workflow

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Log in as a regular user (non-admin) | n/a | n/a | n/a |
| 1 | Create a new case and in the Treatment Plan configure the following: - Configure one implant-borne crown restoration on each arch, i.e. on Upper and Lower Select the Lower pre-op additional scan | 6881_102, 6881_103 | As soon as the implant is configured on each arch, the Lower emergence profile and Upper emergence profile additional scans become automatically selected under Additional scans | Observation/ Screenshot |
| 2 | Unselect the Lower emergence profile additional scan | 6881_104 | Clicking the scan or its checkbox in the Additional scans has no effect, i.e. it remains selected | Observation/ Result Documentation |
| 3 | Click Reset all to remove all treatment plan configurations | 6881_103 | The Lower emergence profile and the Upper emergence profile additional scans become automatically unselected | Observation/ Screenshot |
| 4 | Redo the same configuration as step 1 | n/a | n/a | Preparation |
| 5 | Click Next and observe the different workflow steps displayed under the Scan section of the interface. | 6881_102, 6881_105 6881_201 | The following scan steps are displayed: Upper emergence profile, Lower pre-op, Bite, Lower emergence profile, Lower scanbody. Upper scanbody The Lower pre-op and the Upper emergence profile scan steps are before/above the Bite step, and the rest are after/below the Bite step. Every scan step after/below the Bite is disabled (grayed out) | Observation/ Screenshot |
| 6 | Click on the Lower emergence profile, Lower scanbody, Upper scanbody scan steps in the Scan menu | 6881_201 | Clicking on those steps has no effect, i.e. the steps cannot be opened or navigated to. | Observation/ Result documentation |
| 7 | Perform the Lower pre-op scan and after finishing the scan, inspect the Scan steps menu | 6881_102, 6881_201 | The scan is performed successfully When the scan is completed, the Lower emergence profile step becomes enabled in the Scan menu | Observation/ Screenshot |
| 8 | Perform the Upper emergence profile scan and after finishing the scan, inspect the Scan steps menu | 6881_201 | The scan is performed successfully When the scan is completed, the Upper scanbody step and the Bite step become enabled in the Scan menu | Observation/ Screenshot |
| 9 | Click the Lower emergence profile scan step in the Scan steps menu | 6881_204 | Clicking the step navigates to that step, and: - The ancestor mesh acquired in the Lower pre-op step is displayed | Observation/ Screenshot |
| 10 | Open the HD areas toolkit and add one over a tooth. Go to the next scan step. Then, go back to the first scan. | 6881_308, 6881_313 | HD marker is added over the tooth. After going to the second step and back to the first, the cylinder previously added was saved and it is visible. | Observation/ Screenshot |
| 11 | Click the Lower scanbody scan step | 6881_205 | Clicking the Lower scanbody step navigates to that step, and: - The ancestor mesh acquired in the Lower emergence profile step is displayed, as well as the HD area | Observation/ Screenshot |
| 12 | Click the Upper scanbody scan step | 6881_205 | Clicking the Upper scanbody step navigates to that step, and: - The ancestor mesh acquired in the Upper emergence profile step is displayed | Observation/ Screenshot |

## TC06113 – Complex workflow – Mixed restoration workflow

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Log in as a regular user (non-admin) | n/a | n/a | n/a |
| 1 | Create a new case, and in the Treatment Plan, configure the following: - One implant-borne crown restoration on the Upper arch - Configure also one crown restoration on the upper arch. | 6881_103 | Once the implant is configured on the arch, the Upper emergence profile becomes automatically selected. | Observation/ Screenshot |
| 2 | Select the Upper pre-op additional scan | 6881_102 | Additional pre-op scan can be added to the treatment plan. | Observation/ Screenshot |
| 3 | Unselect the Upper emergence profile additional scan | 6881_104 | Clicking the scan or its checkbox in the Additional scans has no effect, i.e. it remains selected | Observation/ Result Documentation |
| 4 | Click Reset all to remove all treatment plan configurations | 6881_103 | The Upper emergence profile additional scan becomes automatically unselected | Observation/ Screenshot |
| 5 | Redo the same configuration as step 1 | n/a | n/a | Preparation |
| 6 | Click Next and observe the different workflow steps displayed under the Scan section of the interface. Note the presence of scanbody steps. | 6881_105 | The following scan steps are displayed: Lower, Upper emergence profile, Upper pre-op, Bite and Upper scanbody. | Observation/ Screenshot |
| 7 | Check the order of the scan steps and their availability. | 6881_201 | The Upper pre-op scan step is before/above the Bite step, and the rest are after/below the Bite step. Every scan step after/below the Bite is disabled (grayed out) | Observation/ Screenshot |
| 8 | Click on the Upper emergence profile. | 6881_201 | Clicking on this step has no effect, i.e. the steps cannot be opened or navigated to. | Observation/ Result documentation |
| 9 | Click on the Upper scanbody scan step. | 6881_201 | Clicking on this step has no effect. | Observation/ Result documentation |
| 10 | Perform the Upper pre-op scan and after finishing the scan, inspect the Scan steps menu | 6881_102, 6881_201 | The scan is performed successfully When the scan is completed, the Upper emergence profile step becomes enabled in the Scan menu | Observation/ Screenshot |
| 11 | Perform the Lower, after finishing the scan, inspect the Scan steps menu | 6881_201 | The scan is performed successfully When the scan is completed, the Bite step becomes enabled in the Scan menu | Observation/ Screenshot |
| 12 | Click the Upper emergence profile scan step in the Scan steps menu | 6881_204 | The ancestor mesh acquired in the Upper pre-op step is displayed The Upper scanbody step become enabled in the Scan menu. | Observation/ Screenshot |
| 13 | Perform the Upper emergence profile scan, after finishing the scan, inspect the Scan steps menu | 6881_201 | The scan is performed successfully | Observation/ Screenshot |
| 14 | Open the HD areas toolkit and add one over a tooth. Go to the next scan step. Then, go back to the first scan. | 6881_308, 6881_313 | HD marker is added over the tooth. After going to the second step and back to the first, the cylinder previously added was saved and it is visible. | Observation/ Screenshot |
| 15 | Click the Upper scanbody scan step | 6881_205, 6881_301 | The ancestor mesh acquired in the Upper emergence profile step is displayed, as well as the HD area | Observation/ Screenshot |

## TC06119 – HD areas – Markers and toolkit

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Use arch model with tooth preparation Create basic and crown and bridge workflows, add implants and crowns with pre-ops Perform all the scans in all the cases Proceed with the following steps in the two cases Availability of the HD areas toolkit will be checked only in the scan steps in which it is shown (all except pre-op and bite steps) | n/a | n/a | n/a |
| 1 | Check the HD areas availability in all the scan steps. | 6881_301, 6881_306 | The HD areas toolkit is available and closed by default. | Observation/ Screenshot |
| 2 | Move from one scan step to another without using the HD areas toolkit. | 6881_301, 6881_305 | User moves to other scan step without interacting with the HD toolkit. | Observation / Result documentation |
| 3 | Open the HD areas toolkit in a crown and bridge case scan step with a tooth preparation and a scanbody surfaces. | 6881_310 | HD area markers are automatically detected and placed only for the mentioned setup. | Observation/ Screenshot |
| 4 | Move to a derived scan having the ancestor HD areas marked. | 6881_313 | HD area markers placed on an ancestor scan are included here. | Observation/ Screenshot |
| 5 | Open the HD areas toolkit and click on a tooth. Then, click on other teeth to Add several HD area markers in one of the steps, i.e. Lower. Go to another scan step and go back to the previous one. | 6881_308 | HD markers added are still displayed. | Observation/ Screenshot |
| 6 | Open the Fairy mother tool (Ctrl+Alt+S and then, Ctrl+Alt+F1). Go to SharedScene tab and go to Appearance section of the arch, Polygon Mode, then choose Point or Line and check the HD markers. | 6881_312 | HD markers are identified on the mesh. | Observation/ Screenshot |
| 7 | Rotate the 3D mesh and open the HD areas toolkit again. | 6881_306 | The 3D mesh is reoriented automatically and positioned in top view. | Observation/ Screenshot |
| 8 | Rotate the arch 3D mesh. | 6881_308 | HD markers keep their orientation with the 3D mesh. HD markers cylinders rote with the mesh. | Observation/ Screenshot |
| 9 | Open the HD areas toolkit again and move the diameter handle of any cylindrical shape in any direction. | 6881_303 | The diameter of the area of interest increases/decreases following the outward/inward movements of the diameter handle. | Observation/ Screenshot |
| 10 | Adjust the size of the HD markers moving the pointer in the cylinders. Move one of the cylinders by clicking and dropping it over another tooth. | 6881_303 | - The size of the HD markers is adjusted; user makes it greater or lower. - The HD marker is moved to another tooth. | Observation/ Screenshot |
| 11 | Try to move one of the cylinders outside the perimeter of the mesh. | 6881_309 | HD marker cannot be moved outside the perimeter of the mesh. | Observation / Result documentation |
| 12 | Close the HD area toolkit and check the mesh. | 6881_307 | HD markers are not displayed when the toolkit is closed. | Observation/ Screenshot |
| 13 | Open the HD areas toolkit, click on Remove option and click on one of the HD markers in the mesh. Remove several HD markers. | 6881_304 | HD markers are removed, and the cylinders are not shown anymore. | Observation/ Screenshot |
| 14 | Close and reopen the case. Navigate between the scan views with HD marked areas. | 6881_306, 6881_307, 6881_308 | - Case is properly reopened. - HD area markers are properly displayed. | Observation / Result documentation |
| 15 | Proceed to Review the case. Go to a scan step and remove an HD area marker. Then, check the Review view. | 6881_311 | - HD marker is removed. - Review is not available. | Observation / Result documentation |

## TC06114 – Complex workflow – Ancestry conflicts – Selecting/unselecting pre-op scan

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create a new case, and in the Treatment plan, configure the following: - - Implant restoration on one tooth in the Upper arch - Select the Upper pre-op additional scan Acquire the Lower scan, the Upper pre-op scan and the Upper emergence profile scan then start the test from the interface of the Upper emergence profile scan step. | n/a | n/a | n/a |
| 1 | Click Delete, then click Yes | 6881_202 | When clicking Delete: - A warning dialog appears on top of the interface and explains that scan data will be lost When clicking Yes: - The mesh is displayed automatically from the top, i.e. the occlusal surface | Observation/ Screenshot |
| 2 | Click on the Upper scanbody scan step in the Scan menu | 6881_202 | Scanbody step is accessible from the Scan menu | Observation/ Screenshot |
| 3 | Navigate back to the Upper emergence profile scan step and perform the scan again | n/a | n/a | Preparation |
| 4 | Navigate to Treatment plan interface and uncheck the Upper pre-op scan under Additional scans | 6881_202 | - A warning dialog appears on top of the interface and explains that scan data will be lost | Observation/ Screenshot |
| 5 | Confirm the deletion of data in the warning dialog. | 6881_202 | The Upper pre-op scan’s checkbox becomes unchecked | Observation/ Screenshot |
| 6 | Save the treatment plan by clicking Next. | n/a | n/a | Preparation |
| 7 | Click Next and observe the scan steps displayed under the Scan menu | 6881_202 | The following scan steps are now displayed: Lower, Upper emergence profile, Bite, Upper scanbody | Observation/ Screenshot |
| 8 | Navigate to the Upper emergence profile scan step | 6881_202 | -There is no more scan data in that scan step, i.e. it is empty - The title(header) of the interface for that scan step is now “Upper emergence profile” | Observation/ Screenshot |
| 9 | Navigate to the Treatment plan interface, then successively check and uncheck the Upper pre-op scan under Additional scans | 6881_202 | No warning dialog is displayed | Observation/ Screenshot |
| 10 | Click on the Upper emergence profile scan step and perform that scan. | n/a | n/a | Preparation |
| 11 | Navigate to the Treatment plan interface and check the Upper pre-op scan under Additional scans | 6881_202 | A warning dialog appears on top of the interface and explains that scan data will be lost | Observation/ Screenshot |
| 12 | Confirm the deletion of data in the warning dialog. | 6881_202 | The Upper pre-op scan’s checkbox becomes checked | Observation/ Screenshot |
| 13 | Save the treatment plan by clicking Next. | n/a | n/a | Preparation |
| 14 | Click Next and observe the scan steps displayed under the Scan menu | 6881_202 | The following scan steps are now displayed: Lower, Upper pre-op, Bite, Upper emergence profile, Upper scanbody | Observation/ Screenshot |
| 15 | Navigate to the Upper emergence profile scan step Navigate to the Upper scanbody scan step by clicking on it | 6881_201, 6881_202 | - The steps are grayed out in the Scan menu - Clicking on each of these steps has no effect, i.e. the step cannot be opened | Observation/ Screenshot |
| 16 | Navigate to the Upper pre-op scan step and perform that scan | n/a | n/a | Preparation |
| 17 | Navigate to the Upper emergence profile scan step | 6881_201, 6881_202 | The mesh of the Upper pre-op step is preloaded and intact | Observation/ Screenshot |
| 18 | Navigate to the Lower scan step | 6881_202 | The mesh acquired in Setup is unaffected in this scan step. | Observation/ Result documentation |

## TC06115 – Complex workflow – Ancestry conflicts - Selecting/unselecting implant in treatment plan

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create a new case, and in the Treatment plan, configure the following: - - Implant crown restoration on two teeth in the Upper arch - - Select the Upper pre-op additional scan Acquire the Upper pre-op scan, the Upper emergence profile scan and the Upper scanbody scan | n/a | n/a | n/a |
| 1 | Navigate to the Treatment plan interface and remove one of the two implant configurations from its tooth on the Upper arch, by using the Clear button under the Implant kit menu. | 6881_202 | The implant part is removed from the tooth’s crown configuration. | Observation/ Screenshot |
| 2 | Repeat step 1, this time for the second (and last) tooth on the Upper arch. | 6881_202 | A dialog is shown over the treatment plan, indicating that scan data will be lost if the user proceeds with this action. | Observation/ Screenshot |
| 3 | Cancel the deletion of data in the warning dialog, i.e. refuse the confirmation. | 6881_202 | After the cancellation, the dialog closes and the implant configuration remains on the tooth. | Observation/ Result documentation |
| 4 | Repeat step 2 and 3 but this time confirm the deletion of data in the warning dialog that appears. | 6881_202 | The implant part is removed from the tooth’s crown configuration. | Observation/ Result documentation |
| 5 | Proceed to the Scan interface and navigate to the Upper pre-op scan step | 6881_202 | Scan data acquired during Setup in this step is unaffected, i.e. the scan is still there. | Observation/ Screenshot |
| 6 | Navigate to Treatment Plan and click Reset all | 6881_202 | The entire treatment plan’s dental chart is void of any configuration. | Observation/ Screenshot |
| 7 | Navigate to the Upper scan step by clicking on it | 6881_202 | When opening the step, only the scan mesh acquired in Upper pre-op is displayed with no other scan data. | Observation/ Screenshot |
| 8 | Perform the Upper arch scan step, then navigate to the Treatment plan interface and configure one implant on the Upper arch | 6881_202 | When the implant configuration is completed: A warning dialog appears on top of the interface and explains that scan data will be lost. | Observation/ Screenshot |
| 9 | Confirm the deletion of data in the warning dialog and Save the treatment plan | 6881_202 | The implant configuration is applied to the tooth and the treatment plan is saved successfully. | Observation/ Result documentation |
| 10 | Proceed to the Scan interface and navigate to the Upper emergence profile scan step | 6881_202 | The mesh of the Upper pre-op step is preloaded and intact | Observation/ Screenshot |
| 11 | Navigate to the Upper scanbody scan step by clicking on it | 6881_201 | Scanbody step is accessible from the Scan menu | Observation/ Screenshot |

## TC06116 – Complex workflow – Ancestry conflicts - Swapping implant configurations in treatment plan

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create a new case, and in the Treatment plan, configure the following: - - Implant crown restoration on two teeth in the Upper arch - - Select the Upper pre-op additional scan Acquire all scans in the workflow | n/a | n/a | n/a |
| 1 | Navigate to the Treatment plan interface, select one of the two teeth configurations of the Upper arch and move it to any tooth on the Lower arch, by using the Swap button. | 6881_202 | A dialog is shown over the treatment plan, indicating that scan data will be lost if the user proceeds with this action. | Observation/ Screenshot |
| 2 | Confirm the deletion of data in the warning dialog and observe the treatment plan interface. | 6881_202 | The tooth configuration is moved to the Lower arch. | Observation/ Screenshot |
| 3 | Proceed to the Scan interface and inspect every scan step for the Lower arch. | 6881_202 | Every scan step for the Lower arch is void of scan data, i.e. empty. | Observation/ Screenshot |
| 4 | Perform the scan for the Lower emergence profile step | n/a | n/a | Preparation |
| 5 | Repeat step 1, this time for the second (and last) tooth on the Upper arch. | 6881_202 | A dialog is shown over the treatment plan, indicating that scan data will be lost if the user proceeds with this action. | Observation/ Result documentation |
| 6 | Confirm the deletion of data in the warning dialog and observe the treatment plan interface. | 6881_202 | The tooth configuration is moved to the Lower arch. The Upper arch is now void of any configuration. | Observation/ Screenshot |
| 7 | Proceed to the Scan interface and inspect every scan step for the Upper arch. | 6881_202 | The scan of the Upper pre-op step is still present and unaffected. The Upper scan step contains a preloaded Upper pre-op scan with no other scan data. | Observation/ Result documentation |
| 8 | Proceed to the Scan interface and inspect the Lower emergence profile step. | 6881_202 | The scan acquired previously for the Lower emergence profile step is intact. | Observation/ Result documentation |

## TC06117 – Scan deletion – No treatment plan checkbox

| Step | Description | Requirement | Expected Result | Test Method / Objective Evidence |
| --- | --- | --- | --- | --- |
| Setup | Create a new Quick Scan, then acquire all scans. | n/a | n/a | n/a |
| 1 | Navigate to Treatment plan and check the No treatment plan checkbox. | 6881_202 | The No treatment plan checkbox is checked. | Observation/ Result documentation |
| 2 | Remove the “Quick Scan” patient, add a new patient, and click Next, then inspect each scan step in the Scan interface. | 6881_202 | The scan data of every scan step is unaltered, i.e. identical as when it was scanned during Setup step. | Observation/ Result documentation |
| 3 | Navigate to Home and create a New Case with at least one implant configuration on a single arch. Acquire all scans in the workflow, then navigate to the Treatment plan and check the No treatment plan checkbox. | 6881_202 | A dialog is shown over the treatment plan, indicating that scan data will be lost if the user proceeds with this action. | Observation/ Screenshot |
| 4 | Confirm the deletion of data in the warning dialog and observe the treatment plan interface. | 6881_202 | The No treatment plan checkbox is checked. The dental chart is void of any tooth configuration. | Observation/ Screenshot |
| 5 | Proceed to the Scan interface and inspect every scan step corresponding to the arch where the implant was configured previously. | 6881_202 | Every scan step for the treated arch is empty of scan data. | Observation/ Result documentation |
| 6 | Inspect every scan step for the opposite arch. | 6881_202 | Every scan step of the opposite arch is intact. | Observation/ Result documentation |
| 7 | Navigate to Home and create a New Case with one tooth-borne restoration on a single arch (without any additional scan). Acquire all scans in the workflow, then navigate to the Treatment plan and check the No treatment plan checkbox. | 6881_202 | The No treatment plan checkbox is checked. The dental chart is void of any tooth configuration. | Observation/ Screenshot |
| 8 | Save and click Next, then inspect each scan step in the Scan interface. | 6881_202 | The scan data of every scan step is unaltered. | Observation/ Result documentation |

# Referenced Documents

- DEV-0044563 STMN IOS SDD - Scan Workflows Rev 5

- DEV-0044555 STMN IOS Software Product Requirements Specification Rev 15

- DEV-0044595 STMN IOS Scan Workflow Verification Report

# Anomalies

If anomalies other than the ones included in this specification occur, a summary description is given in the related test report, in Verification and Validation Summary Report (FMR-000716).

# Document History

| Version | Author | Change to previous version |
| --- | --- | --- |
| 1.0 | Habone Youssouf Vincent Bolduc | Initial version |
| 1.1 | Vincent Bolduc | TC06101: Clarified the Setup step instructions TC06103 & TC06110: Added new steps to test SDD 6881_037 TC06104 & TC06105: Updated to test that removed marks are not saved TC06007: Minor vocabulary corrections TC06115: Updated to cover updated design as per SDD 6881_007 Added new TC06116, TC06117 to cover updated design as per SDD 6881_007 Reason: 1.0.0 test deviations/recommendations and changes for release 1.0.1 |
| 3 | Vincent Bolduc | Corrective updates as per Redmine #67582 and #67535: - TC06112 step 8: Added the Bite step as enabled, in the expected result - TC06117 Step 2: Updated to be consistent with the Setup step and avoid unplanned extra operations from the tester - TC06117 step 3: Missing word “least” added in step description. Update TC06110 steps 1, 13, 16: Removed the previous Shine Control tool and renamed the “Soft tissue removal” tool to “Soft tissue level” Updated version number (this section) and items under section Referenced Documents to comply with ETQ transfer Reason: 1.0.1 test deviations/recommendations and changes for STMN IOS release 1.0.2 |
| 4 | Habone Youssouf | Corrective updates as per Redmine #68723: Added TC06108 step 4. Added new coDiagnostiX test TC06118 Reason: 1.0.2 test deviations/recommendations and changes for STMN IOS release 1.0.3 |
| 5 | Teresa Toledano | Requirement Ids renamed as per SDD in: TC06101: 6881_001 → 6881_101 TC06102: 6881_012 → 6881_301 TC06106: 6881_034 → 6881_106 TC06107: 6881_034 → 6881_106 6881_035 → 6881_107 6881_036 → 6881_108 TC06108: 6881_034 → 6881_106 6881_035 → 6881_107 6881_036 → 6881_108 TC06109: 6881_034 → 6881_106 6881_035 → 6881_107 6881_036 → 6881_108 TC06118: 6881_001 → 6881_101 TC06111: 6881_002 → 6881_102 6881_006 → 6881_201 6881_008 → 6881_203 TC06112: 6881_002 → 6881_102 6881_003 → 6881_103 6881_005 → 6881_105 6881_006 → 6881_201 6881_009 → 6881_204 6881_010 → 6881_205 TC06114: 6881_006 → 6881_201 6881_007 → 6881_202 TC06115: 6881_006 → 6881_201 6881_007 → 6881_202 TC06116: 6881_007 → 6881_202 TC06117: 6881_007 → 6881_202 Replace Simple workflow by Basic workflow in TC06101, TC06102, TC06103, TC06104 and TC06107 to avoid confusion with new Simple Scan workflow 6881_011, 6881_020, 6881_021, 6881_022, 6881_023, 6881_024, 6881_025, 6881_026, 6881_027, 6881_028, 6881_029, 6881_030, 6881_031, 6881_032 and 6881_033 removed from SDD: Cut and scan test steps removed in TC06108 and TC06109 TC06103, TC06104, TC06105, TC06106 and TC06110 removed and TC06111 and TC06112 updated: they tested mark and cut feature. TC06113 updated: cut and scan features removed and test updated to cover mixed restoration workflow in 6881_204 and 6881_205 of SDD. Updated to cover 6881_308 and 6881_313 TC06111 and TC06112 updated to cover 6881_206, 6881_308 and 6881_313 TC06119 created and TC06102 updated to test HD area markers and toolkit and HD approach reminder: 6881_301, 6881_302, 6881_303, 6881_304, 6881_305, 6881_306, 6881_307, 6881_308, 6881_309, 6881_310, 6881_311, 6881_312 and 6881_313, 6881_109 and 6881_110 Added 6881_206, 6881_308, 6881_313 TC06118 updated as per Simple Scan, not coDiagnostiX anymore, and to cover 6881_308 Added Model-specific Test cases section Reason: Release 1.0.7 (DEV-0044548 Rev 8) |
| 6 | Martin Carufel Shivani Girish Munale | Modify TC06111: Move step 6 to step 5, and split previously step 5 (now step 6) into step 6 and 7 Modify TC06113: Insert new step (11) after step 10. Move step 12 the newly created step 11. Step 12 change the expected result. Step 13 change expected result. Modify TC06114: minor changes to steps 2, 6, 9, 13, and expected result of step 1 and 17 Modify TC06119: Setup add a specific model type. Modifiy TC06115 steps 11 and expected result of step 10 and 11. Reason: Release STMN IOS 1.0.7 release (DEV-0044548 Rev 8) |

# Review & Approval

Content Approver of this document confirms that the Reviewer(s) listed below have reviewed and approved the content of this document version.

| Project Role | Function | Name | Date |
| --- | --- | --- | --- |
| Software Quality Analyst | Author | Martin Carufel | 2026-03-31 |
| Software Quality Analyst | Reviewer | Teresa Toledano | 2026-03-31 |
| Quality and Regulatory Expert | Reviewer | Rama Almaleh | Electronic Signature |
| SQA SME | Approver | Natalia Piworun | Electronic Signature |
| DEV SME | Approver | Jean-David Gaudreau | Electronic Signature |

# Signatures

Electronic Signature